## Slide 1: Visual Setup Guide
PABV Ventilator
Version 1.0

## Slide 2: Table Of Contents
Unpacking and taking inventory
Setup Pneumatics
  From regulated “hospital air” or compressor
  From high pressure cylinder bottle
Setup Patient Circuit
Setup Electronics

## Slide 3: 1) Unpacking and Taking Inventory

## Slide 4
Input Pneumatic Air Supply
Needs 30-60 PSI
Choose the most appropriate barb to adapt your air supply to push-to-connect
Run Polyflow tube to RED input valve
Push to connect coupler
Adapt barb to Polyflow ¼” tube
Note: Cut Polyflow with razor blade against a hard surface
Tubing
Release Ring
Note: Make connection directly to wall LAST, do not use a flow regulator
2.1) Connect pneumatic air supply to hospital air

## Slide 5
Input Pneumatic Air Supply
Needs 40-60 PSI
Run Polyflow tube to RED input valve
Note: Cut Polyflow with razor blade against a hard surface
2.2) Connect pneumatic air supply to other
High pressure gas cylinders with regulator output, set to 50 psi.
Compressor output with sufficient capacity has regulated pressure output, set to 50 psi.
Note: A failsafe backup is recommended and can be valved in using one-way check valves
And/or

## Slide 6: 3) Patient Circuit
Spirometer Sense Tubes
(to sensor box)
HME Filter
To Patient
Spirometer
PEEP
HEPA Filter
Adaptor
MRI-Conditional Transport Ventilator Circuit
PIP Valve

## Slide 7: 4) Setup Electronics
Computer Data Cable
Sensor Box cable
Pneumatics cable